# Orascan Project: Comprehensive Technical Analysis Report

Executive Summary

Orascan is an innovative non-invasive oral health screening device designed to democratize access to dental healthcare through advanced automation and AI-powered diagnostics. The project has successfully developed a functional prototype that combines multi-modal data collection (imaging, audio, biochemical sensing) with intelligent positioning systems and automated report generation.

#### Key Achievements:

✅ Complete 13-position oral imaging system with automated positioning

✅ Multi-modal data collection: images, TMJ audio, H2S gas monitoring

✅ AI-powered diagnostic reports using YOLO object detection

✅ User-friendly touch interface with patient management system

✅ Professional report generation (HTML/PDF) with automated distribution capabilities

#### Target Populations:

Portable Model: Underserved regions, pregnant women screening programs, dental camps in remote areas

Kiosk Model: Educational institutions, corporate offices, high-traffic areas in tier-one cities

Current Status: MVP ready for pilot testing with identified improvements needed before production deployment.

# Project Background

## Purpose and Vision

Orascan addresses the critical gap in oral healthcare accessibility by providing a non-invasive, automated screening solution that can be operated by trained technicians rather than requiring specialized dental professionals. The system aims to enable early detection of oral diseases, reduce screening costs by 70-80%, and increase screening capacity by 10x through remote dental consultation capabilities.

## Target Audience

### Portable Model Applications:

- Underserved Regions: Rural areas lacking specialized dental services

- Pregnant Women Programs: Specialized oral health monitoring during pregnancy

- Dental Screening Camps: Mass screening initiatives in remote locations

- Primary Health Centers: Integration with existing healthcare infrastructure

### Kiosk Model Applications:

- Educational Institutions: Routine screening for students and staff

- Corporate Offices: Employee health and wellness programs

- Urban Centers: Convenient access points in high-traffic areas

- Healthcare Facilities: Triage and initial screening capabilities

## Medical Objectives

### Primary Detection Capabilities:

- Tooth Cavities: Current detection threshold: 2mm x 2mm lesions, target: 1mm x 1mm

- Gum Problems: Swelling, bleeding, discoloration, tartar deposits

- Missing Teeth: Gap detection with 4-10mm minimum threshold

- Irregular Teeth: Rotation, tilting, misalignment assessment

- Pus/Swelling: Infection indicators through color and texture analysis

- Abnormal Growths: Early oral cancer detection capabilities

### Advanced Diagnostic Features:

- TMJ Disorder Detection: Audio analysis of joint sounds (clicks, pops, crepitus)

- Biochemical Analysis: H2S gas detection for diabetes indicators and oral health correlation

- Temperature Analysis: Planned IR sensor integration for inflammation detection

# System Architecture

## Hardware Components

### Power System:

- Primary Supply: Suswe 12V 6A SMPS (72W capacity)

- Motor Control: Waveshare Stepper Motor HAT (B) with integrated HR8825 drivers

- System Power: 5V regulation for Raspberry Pi via HAT's MP1584 regulator

- ⚠ Power Issue: Total calculated consumption (79.5W) exceeds supply capacity

### Control System:

- Main Controller: Raspberry Pi 5 (powered through HAT GPIO pins)

- Display: 10.1" Capacitive Touch Display (1280×800 resolution, 10-point multitouch)

- Interfaces: HDMI (video), USB-C (touch), I2C (sensors), UART (future expansion)

### Motion Control System:

#### Stepper Motors (Precision Positioning):

- 2x NEMA 17 Motors (200 steps/revolution, 1.8° per step, 2A per phase)

- X-axis Control: GPIO 24 (DIR), GPIO 18 (STEP), GPIO 23 (EN)

- Y-axis Control: GPIO 13 (DIR), GPIO 19 (STEP), GPIO 12 (EN)

- Microstepping: Configurable (full, 1/2, 1/4, 1/8, 1/16, 1/32 steps)

#### Servo Motors (Camera Positioning):

- 2x MG90S Servos (180° rotation range, 1.8-2.2 kg/cm torque)

- Pan Control: GPIO 27 (X-axis camera movement)

- Tilt Control: GPIO 17 (Y-axis camera movement)

- PWM Control: 50Hz frequency, 1-2ms pulse width range

Sensing Systems:

| Sensor Type | Model/Interface | Purpose | Status |
| --- | --- | --- | --- |
| Imaging | USB Camera (CV2) | Oral cavity imaging | ✅ Implemented |
| Gas Detection | MQ316 via ADS1115 I2C | H2S concentration monitoring | ✅ Implemented |
| Audio | USB Microphone | TMJ sound recording | ✅ Implemented |
| Temperature | IR Array (MLX90640) | Inflammation detection | ⚠ Planned |

GPIO Pin Mapping:

| GPIO Pin | Physical Pin | Function | Component |
| --- | --- | --- | --- |
| GPIO 24 | Pin 18 | X-axis Direction | Stepper Motor HAT |
| GPIO 18 | Pin 12 | X-axis Step | Stepper Motor HAT |
| GPIO 13 | Pin 33 | Y-axis Direction | Stepper Motor HAT |
| GPIO 19 | Pin 35 | Y-axis Step | Stepper Motor HAT |
| GPIO 27 | Pin 13 | Pan Servo PWM | Camera Pan Control |
| GPIO 17 | Pin 11 | Tilt Servo PWM | Camera Tilt Control |
| I2C SDA/SCL | Pins 3,5 | Sensor Communication | ADS1115 (H2S Sensor) |

## Software Architecture

### Technology Stack:

- Platform: Raspberry Pi OS (Debian-based Linux)

- Programming Language: Python 3.x

- GUI Framework: Tkinter with ttk styling (responsive design)

- Database: MySQL with PyMySQL/mysql-connector drivers

- Computer Vision: OpenCV (cv2), YOLO via ultralytics

- Hardware Control: lgpio library (modern GPIO interface)

- Data Visualization: Matplotlib with Tkinter backend

- Audio Processing: PyAudio, wave libraries

- Sensor Interface: Adafruit CircuitPython (ADS1115 I2C)

- Report Generation: FPDF (PDF), HTML with base64 image embedding

### System Modules:

- login.py: User authentication, patient dashboard, profile management

- automatic.py: Automated image capture with predefined positioning

- manual.py: Manual image capture with user-controlled positioning

- audio_data.py: TMJ audio recording with waveform visualization

- breath_data.py: H2S gas monitoring with peak detection

- HR8825.py: Stepper motor driver implementation

- lgpio.py: GPIO control wrapper functions

- ReportGenerator class: YOLO-based analysis and report generation

# Work Completed So Far

  - User	Management	System	✅	Fully Implemented

### Authentication & Registration:

Patient registration with comprehensive validation (email format, 10-digit phone, Aadhar number)

Secure login system with email/password authentication Password reset functionality with email verification Session management (basic implementation)

### Patient Dashboard:

Home Section: Patient profile display with personal information Submit Data: Access to image, audio, and breath data collection Edit Profile: Update personal information, contact details

Help Section: Contact information and support resources

Database Schema:

patient (id, f_name, l_name, contact, email, place, Aadhar, password)

  - Image	Capture	System	✅	Fully Implemented

### Manual Mode Features:

13 Standardized Positions: Complete oral cavity coverage with predefined angles

Visual Guidance: Reference images and animated GIF demonstrations for each position

Step-by-Step Instructions: Detailed text instructions for patient positioning

Live Camera Feed: Real-time preview with positioning controls

#### Manual Controls:

Servo Control: Pan/tilt for camera positioning (±10° increments) Stepper Control: X/Y axis movement (100-step increments) Reset function: Return to center position

Image Preview: Side-by-side comparison with reference image

Capture Options: Retake, Save, Save & Next workflow

Responsive UI: Fullscreen support, window resizing, touch-optimized

### Automatic Mode Features:

Predefined Positioning: SERVO_POSITIONS array with 13 optimized camera angles

Timer-Based Capture: 2-second intervals between positions

Auto-Progression: Automatic advancement through all 13 positions Visual Feedback: Screen flash confirmation on successful capture Safety Features: Auto-tracking disabled during capture sequences

### Technical Implementation Details:

- Camera Integration: cv2.VideoCapture with USB camera support

- Storage System: Dual storage (MySQL BLOB + local filesystem)

- Image Identification: Unique identifiers (phone_date_sequence format)

- Quality: Full resolution JPEG with configurable compression

- Threading: Thread-safe camera operations for UI responsiveness

  - Audio Capture System ✅ Fully Implemented

### TMJ Audio Recording:

Duration: Fixed 10-second recording sessions

Quality: 44.1kHz sample rate, 16-bit PCM, mono channel

Format: WAV file output for maximum compatibility

Real-time Visualization: Live waveform display during recording

Storage: MySQL database (BLOB) + local file backup

Patient Association: Automatic linking to patient ID

Multi-threading: Non-blocking recording for responsive UI

Device Detection: Automatic USB microphone enumeration

### Database Schema:

audio_captures (patient_id, audio_data, timestamp)

  - H2S Gas Monitoring ✅ Fully Implemented

### Sensor Configuration:

Sensor Model: MQ316 H2S sensor via ADS1115 I2C ADC

Calibration: R0 = 27.12fi baseline in clean air

Noise Compensation: 0.72 ppm baseline subtraction

Measurement Range: 0-50 ppm with 0.1 ppm resolution

### Data Processing:

Windowed Analysis: 10-12 second measurement windows

Peak Detection: Maximum value extraction per window

Real-time Display: Live graph visualization with threshold lines

Alert System: Configurable threshold warnings

Medical Relevance: Diabetes correlation indicators, oral health assessment

Database Schema:

h2s_peaks (patient_id, peak_concentration, timestamp)

  - Motor Control System ✅ Fully Implemented

### Stepper Motor Implementation:

Driver: HR8825 with custom Python class implementation

Microstepping: Configurable resolution (full, 1/2, 1/4, 1/8, 1/16, 1/32 steps)

Control Modes: Software and hardware microstepping selection Directional Control: Bidirectional movement (forward/backward) Speed Control: Adjustable step delay (default: 0.005s) Configuration:

X-axis: Full-step mode (software control) Y-axis: Half-step mode (hardware control)

### Servo Motor Implementation:

Custom StableServo Class: Anti-jitter implementation

Angle Range: -90° to +90° (or 0° to 180° depending on configuration)

PWM Control: 500-2500µs pulse width, 50Hz frequency

Smart Control: Auto-detach after positioning to prevent jitter Positioning Modes: Manual control and automatic predefined positions Safety Features: Center reset function, position memory

  - Report	Generation	System	✅	Fully

## Implemented

### AI-Powered Analysis:

YOLO Integration: Pre-trained model (/home/pi/Downloads/best.pt) Batch Processing: Automatic analysis of all 13 captured images Detection Classes: Cavity, lesion, plaque with confidence scoring Severity Classification: High/medium/low risk categorization

### Report Generation:

HTML Reports: Professional layout with embedded base64 images

PDF Export: FPDF-based PDF generation capability

Image Comparison: Side-by-side original vs. detection visualization

Findings Summary: Structured diagnostic output with confidence levels Distribution Ready: Email integration with SMTP and attachments Web Serving: Flask server on port 5000 for report access

Cleanup: Automatic temporary file management

-- Core Tables Implemented patient ( id INT PRIMARY KEY AUTO_INCREMENT, f_name VARCHAR(50), l_name VARCHAR(50),

contact VARCHAR(10), email VARCHAR(100) UNIQUE,

place	VARCHAR(100),	Aadhar	VARCHAR(12),	password

VARCHAR(255) );

patient_images ( INT, image_data LONGBLOB, created_at

TIMESTAMP,

patient_id INT, image_number

image_identifier VARCHAR(50),

file_path

VARCHAR(255)

);

audio_captures ( patient_id INT, audio_data LONGBLOB,

Database	Implementation	✅	Complete Schema

timestamp TIMESTAMP );

h2s_peaks ( patient_id INT, peak_concentration DECIMAL(5,2), timestamp TIMESTAMP );

# Current Capabilities

## What the System Can Do:

- ✅	Complete Patient Management: Registration, authentication, profile management

- ✅ Standardized Image Capture: 13 positions with manual or automatic modes

- ✅	Real-time Guidance: Reference images, animations, and step-by-step instructions

- ✅ TMJ Audio Recording: 10-second high-quality audio capture

- ✅ H2S Gas Monitoring: Real-time concentration measurement with peak detection

- ✅ Precise Camera Positioning: Servo and stepper motor control

- ✅ Comprehensive Data Storage: MySQL database with file system backup

- ✅ AI-Powered Diagnostics: YOLO-based disease detection

- ✅ Professional Reporting: HTML and PDF export capabilities

- ✅ Touch-Optimized Interface: Fullscreen, responsive design

System Performance Metrics:

| Operation | Current Performance | Target Performance |
| --- | --- | --- |
| Image capture (auto mode) | ~2 seconds per image | <1 second per image |
| Complete screening (13 images) | 2-3 minutes | 1-2 minutes |
| Audio capture | 10 seconds (fixed) | 5-10 seconds (configurable) |
| H2S measurement window | 10-12 seconds | 5-10 seconds |

| Report generation | 30-60 seconds | <30 seconds |
| --- | --- | --- |
| Total screening time | 4-5 minutes | 2-3 minutes |

Current Issues:

Code Duplication: Significant overlap between automatic.py and manual.py (estimated 60% shared code)

Hardcoded Paths: File paths hardcoded to /home/pi/Downloads/...

GPIO Inconsistencies: Different pin definitions across modules

Configuration	Management:	Database	credentials	and	settings	scattered throughout code

Error Handling: Inconsistent exception handling across modules

# Areas for Improvement

  - Power Management ⚠ Critical Issue

Current Problem:

Power Deficit: Total calculated power consumption (79.5W) exceeds SMPS capacity (72W)

Risk Impact: System brownouts during peak usage, motor stalling, camera feed interruption

Immediate Solutions:

- Hardware Upgrade: Replace with 90W or 100W SMPS (2 days)

- Current Monitoring: Add real-time power consumption tracking

- Power Management: Implement motor sleep modes when inactive

- Load Balancing: Stagger motor operations to reduce peak draw

  - Code Quality and Maintenance 🔧 Technical Debt

Recommended Solutions:

- Refactoring: Create shared base class for image capture functionality

- Configuration Files: JSON/YAML files for paths, credentials, and settings

- GPIO Centralization: Single module for all pin definitions

- Environment Variables: Secure credential management

- Logging Framework: Python logging module implementation

- Error Recovery: Comprehensive try-catch with fallback procedures

  - Database	Optimization	📊	Performance Concerns

### Current Limitations:

BLOB Storage: Large images and audio files stored directly in database Index Strategy: No documented indexing for performance optimization Backup Procedures: No automated backup system

Data Retention: No policy for archiving old records

Connection Management: No connection pooling implementation

### Optimization Plan:

- File System Storage: Move large files to filesystem with database references

- Index Implementation: Add indexes on patient_id, timestamp, email

- Automated Backups: Daily/weekly backup scripts

- Data Archival: Archive records older than 1 year

- Connection Pooling: Implement connection pool for better performance

  - User	Experience	Enhancements	💡	UX Improvements

Progress Indicators:

Missing Features:

Progress bar during report generation Remaining image counter in auto-capture mode Loading animations for database operations

Image Quality Feedback:

Blur detection before saving (Laplacian variance method) Brightness/contrast warnings

Mouth detection confirmation Focus quality assessment

Accessibility & Localization:

Multi-language Support: Currently English only, need Hindi and regional languages

Voice Guidance: Audio instructions for positioning High-Contrast Mode: For visually impaired users Keyboard Navigation: Alternative to touch interface

  - Clinical Validation 🏥 Medical Requirements

### Missing Critical Components:

IR Temperature Sensor: MLX90640 integration not implemented

TMJ Audio Analysis: Only recording implemented, no pattern analysis Model Validation: YOLO model not validated against clinical diagnoses Ground Truth Data: No clinical training dataset

### Required Implementation:

- Temperature Sensing: Integrate IR array for inflammation detection

- Audio Processing: FFT analysis, spectrogram generation for TMJ patterns

- Clinical Validation: Partner with dental professionals for model training

- Dataset Collection: 1000+ annotated images with ground truth diagnoses

- Confidence Thresholds: Implement clinical confidence scoring

6.1.1 Fully Automated Screening Workflow ⭐⭐⭐

Current State: Requires manual button clicks between each image capture

Proposed Enhancement: Complete end-to-end automation

Workflow Design:

Login → Auto Start → 13 Images → Audio → H2S → Report

→ Email/Display

Security	Improvements	🔒	Critical

## Vulnerabilities

Current Security Issues:

Critical: Passwords stored as plain text Critical: Potential SQL injection vulnerabilities High: No session timeout management

High: Missing GDPR/HIPAA compliance features

Medium: No audit logging for medical data access

Critical Security Fixes:

- Password Security: Implement bcrypt/argon2 hashing immediately

- SQL Security: Use parameterized queries throughout (partially implemented)

- Session Management: 15-30-minute timeout, secure session tokens

- Data Encryption: Encrypt medical data at rest

- Audit Trail: Log all data access and modifications

- Compliance: GDPR consent forms, data deletion capabilities

# Automation Opportunities

## High-Priority Automation

6.1.3 Smart Positioning System ⭐⭐⭐

Current Limitation: Fixed predefined positions, no feedback mechanism

Enhancement: Vision-based adaptive positioning

Technical Implementation:

Implementation Requirements:

State Machine: Python state management for workflow progression

Voice Prompts: Audio guidance for patient positioning

Automatic Progression: Timer-based advancement (3-4 minutes total) Error Recovery: Retry logic for failed captures (max 3 attempts) Quality Gates: Automatic quality validation before proceeding

Estimated Development Time: 3 weeks

ROI Impact: Reduce operator involvement by 80%, increase throughput by 3x

6.1.2 Intelligent Image Quality Control ⭐⭐⭐

Current Issue: No quality validation before saving images

Solution: AI-powered quality assessment

Quality Metrics:

Blur Detection: Laplacian variance analysis (threshold: >100) Brightness Check: Histogram analysis for proper exposure Mouth Detection: YOLO/Haar cascade confirmation

View Angle Verification: Ensure proper positioning

Auto-Retry: Maximum 3 attempts per position

Expected Outcome: 95% first-pass image quality, reduce manual review by 70%

Mouth Detection: Real-time detection of mouth location in camera feed

Auto-Centering: Servo adjustment to center mouth in frame

Distance Optimization: Focus adjustment based on mouth size Tracking: Real-time adjustment during capture sequence Calibration: Patient-specific positioning learning

Development Effort: 2 weeks

Benefit: Consistent image quality regardless of patient size/position

6.1.4 Automated Report Distribution ⭐⭐

Current Process: Reports generated but require manual sharing

Automation Goal: Automatic delivery to stakeholders

Distribution Channels:

Email Integration: PDF reports sent to patient and dentist SMS Notifications: Report ready alerts with secure links Portal Integration: Upload to healthcare management systems

Follow-up Automation: Scheduled reminders for high-risk cases

## Medium-Priority Automation

6.2.1 Predictive Maintenance ⭐⭐

Monitoring Systems:

Motor Usage Tracking: Step count, operation hours for stepper motors

Servo Position Errors: Track positioning accuracy degradation

Failure Prediction: Machine learning models for component lifespan

Maintenance Scheduling: Automated maintenance reminders

Parts Management: Automatic reorder notifications

6.2.2 Data Backup and Synchronization ⭐⭐

Automated Systems:

Daily Backups: Automated MySQL database backups

Cloud Sync: AWS S3/Azure Blob storage integration

Offline Queue: Sync pending operations when connectivity restored

Version Control: Report versioning and change tracking

Disaster Recovery: Automated system restore procedures

## Advanced	Automation	(Future Development)

    - AI Diagnostic Enhancements ⭐

Multi-Model Ensemble: YOLO + ResNet + Vision Transformers

Confidence-Based Triage: Automatic urgent/routine classification

Progression Tracking: Compare with previous visits for trend analysis

Risk Scoring: Comprehensive oral health risk algorithms

    - Telemedicine Integration ⭐

Live Consultation: Real-time video consultation mode

Remote Control: Dentist-controlled camera positioning

Real-time Annotation: Live image markup during consultation

E-Prescriptions: Electronic prescription generation

    - Fleet Management ⭐

Multi-Device Dashboard: Centralized monitoring of multiple Orascan units

OTA Updates: Remote software deployment

Usage Analytics: Performance metrics across device fleet

Load Balancing: Patient distribution across available units

# Recommendations	(Prioritized Implementation Plan)

## Phase	1:	Critical	Fixes	(Immediate	-	2 weeks)

Priority: Production Blockers - Must be completed before any deployment

### Power System Upgrade (2 days)

Replace Suswe 12V 6A SMPS with 90W+ unit (10A minimum) Install current monitoring circuit with automatic shutdown

Test under maximum load conditions

Budget: $50-80 for new SMPS + monitoring components

### Security Hardening (5 days)

Implement bcrypt password hashing (replace plain text storage) Fix SQL injection vulnerabilities (complete parameterized queries) Add session timeout management (15-30 minutes)

Basic audit logging for sensitive operations

Effort: 1 senior developer, high priority

- Error Handling & Stability (3 days) Comprehensive try-catch blocks around hardware operations Implement Python logging framework

Create error recovery procedures (camera reconnection, motor reset)

Add system health monitoring

Phase 1 Success Criteria: System stable under load, security vulnerabilities closed, comprehensive error handling

## Phase 2: Quality Improvements (1 month)

Focus: Code quality, maintainability, and basic automation

### Code Refactoring (1 week)

Consolidate automatic.py and manual.py into shared base classes Create configuration management system (JSON/YAML) Centralize GPIO pin definitions

Environment variables for sensitive configuration

- Image Quality Validation (1 week) Implement blur detection (Laplacian variance method) Add brightness/contrast analysis

Mouth detection confirmation before saving

Auto-retry logic (maximum 3 attempts per position)

- Database Optimization (5 days) Move large files (images, audio) to filesystem storage Implement proper database indexing strategy

Set up automated backup system (daily/weekly)

Database connection pooling

### Testing Infrastructure (5 days)

Unit tests for core functions (motor control, image processing) Integration tests for complete workflows

Hardware abstraction layer for testing without Pi Automated test execution framework

Phase	2	Deliverables:	Clean,	maintainable	codebase,	automated	quality validation, robust testing

## Phase	3:	Automation	Implementation	(2 months)

Goal: Transform from semi-manual to fully automated operation

### Fully Automated Workflow (3 weeks)

State machine implementation for complete screening workflow Voice guidance integration (text-to-speech)

Progress indicators and user feedback Error recovery and retry mechanisms

Target: 3-4 minute complete screening with minimal user interaction

### Smart Positioning System (2 weeks)

Computer vision-based mouth detection Auto-centering algorithm for servos Distance optimization based on mouth size Real-time tracking during capture sequence

### Automated Report Distribution (1 week)

Email integration with PDF attachments SMS notification system

Healthcare portal API integration

Scheduled follow-up system for high-risk cases

- Predictive Maintenance (2 weeks) Usage tracking for all mechanical components Failure prediction models

Automated maintenance scheduling

Parts replacement notifications

Phase 3 Outcomes: Fully automated screening, intelligent positioning, proactive maintenance

## Phase 4: Clinical Validation (3 months)

Objective: Medical validation and regulatory preparation

### Missing Sensor Integration (2 weeks)

IR temperature sensor (MLX90640) integration

TMJ audio pattern analysis (FFT, spectrogram generation) Sensor fusion algorithms

Clinical correlation algorithms

- AI Model Enhancement (6 weeks) Clinical dataset collection (1000+ annotated images) YOLO model retraining with medical ground truth

Ensemble model implementation (YOLO + ResNet)

Clinical validation study with dental professionals Confidence threshold optimization

### Localization & Accessibility (2 weeks)

Multi-language support framework

Voice guidance in Hindi and regional languages High-contrast accessibility themes

Keyboard navigation alternatives

### Compliance & Documentation (2 weeks)

GDPR/HIPAA compliance features Medical device documentation Clinical validation reports

User training materials

Phase 4 Milestones: Clinical validation complete, regulatory compliance, multi-language support

## Phase	5:	Advanced	Features	(Ongoing Development)

Future Roadmap: Advanced capabilities for market differentiation

Telemedicine Platform: Live consultation capabilities

Fleet Management: Multi-device monitoring and control

Advanced Analytics: Population health insights

Continuous Learning: Model improvement from usage data

Integration APIs: Healthcare system connectivity

# Technical Debt Summary

## Code Organization Issues:

| Issue Category | Current State | Impact | Estimated Fix Time |
| --- | --- | --- | --- |
| Code Duplication | 8 Python files with 60% overlapping functionality | High maintenance cost, bug propagation | 2 weeks |
| Module Separation | No clear architectural boundaries | Difficult testing, tight coupling | 1 week |
| Configuration Management | Hardcoded paths and credentials throughout | Environment-specific deployment issues | 3 days |
| GPIO Consistency | Pin definitions scattered across files | Hardware conflicts, difficult debugging | 2 days |

Documentation Gaps:

- API Documentation: No formal API specification

- Installation Guide: Missing setup procedures

- Troubleshooting: Limited diagnostic information

- Developer Guide: No onboarding documentation

## Testing Deficiencies:

- Test Coverage: 0% automated test coverage

- CI/CD Pipeline: No automated deployment

- Performance Testing: No benchmarking framework

- Hardware Testing: No mock/simulation capabilities

Technical Debt Clearance Estimate:

Total Effort: 6-8 weeks (1 senior developer + 1 junior developer)

Priority Order: Security fixes → Code refactoring → Documentation → Testing infrastructure

ROI: 50% reduction in bug fixing time, 3x faster feature development

# Conclusion

## Summary of Achievements

The Orascan project has successfully implemented a functional prototype of a comprehensive non-invasive oral health screening system. The achievement represents a significant advancement in accessible healthcare technology, combining multiple data collection modalities with AI-powered diagnostics.

Core Capabilities Successfully Delivered:

✅	Complete Image Capture Pipeline: 13 standardized positions with automated guidance

✅ Multi-Modal Data Collection: Images, TMJ audio, H2S gas concentration

✅	AI-Powered Diagnostic Reports: YOLO-based disease detection with professional reporting

✅ User-Friendly Interface: Touch-optimized, fullscreen operation

✅ Comprehensive Data Management: MySQL database with patient tracking

## Current Development Status: MVP Ready for Pilot Testing

### Key Technical Strengths:

- Comprehensive Data Collection: Multi-sensor approach provides holistic oral health assessment

- Precision Motor Control: Servo and stepper motor integration enables consistent positioning

- AI-Based Disease Detection: YOLO implementation provides automated diagnostic capabilities

- Professional Report Generation: HTML/PDF output suitable for clinical use

- Modular Software Architecture: Separate modules for different functionality areas

Critical Gaps Before Production Deployment

- ⚠ Power System Undersized: 79.5W demand exceeds 72W supply capacity

- ⚠ Security Vulnerabilities: Plain text passwords, potential SQL injection

- ⚠ Clinical Validation Missing: No validation against medical ground truth

- ⚠ Limited Error Handling: Insufficient fault tolerance for production use

- ⚠ Quality Control Absent: No automated image quality validation

Implementation Roadmap

| Timeline | Phase | Key Objectives | Success Criteria |
| --- | --- | --- | --- |
| Week 1-2 | Critical Fixes | Security hardening, power system upgrade | Production-ready stability |
| Month 1-2 | Quality Improvements | Code refactoring, quality validation, testing | Maintainable, reliable codebase |
| Month 3-5 | Automation & Validation | Full automation, clinical validation | Clinically validated, fully automated system |
| 6+ Months | Advanced Features | Telemedicine, fleet management | Market-differentiated capabilities |

Return on Investment Analysis

| Metric | Traditional Screening | Orascan System | Improvement |
| --- | --- | --- | --- |
| Screening Time | 15-20 minutes | 3-4 minutes | 75% reduction |
| Cost per Screening | $25-35 (dentist time) | $5-10 (technician) | 70-80% reduction |

| Specialist Requirement | Licensed dentist | Trained technician | Democratized access |
| --- | --- | --- | --- |
| Daily Capacity | 20-30 patients | 100+ patients | 3-5x increase |
| Remote Consultation | Not possible | 10x dentist efficiency | Scale multiplication |

Strategic Recommendation

Proceed with Phase 1 & 2 improvements immediately before considering wider deployment. The system demonstrates strong technical potential and addresses a critical healthcare need, but requires security hardening, quality improvements, and clinical validation before production use in healthcare settings.

Investment Priority:

- Immediate (2 weeks, $10K): Critical fixes for pilot deployment

- Short-term (2 months, $30K): Quality improvements and basic automation

- Medium-term (6 months, $75K): Clinical validation and advanced automation

- Long-term: Market expansion and advanced features

Market Opportunity